Original Research Article

Agreement and Systematic Differences between Analog and Digital Tools for BMI Assessment in Undergraduate Students in Nigeria

# ABSTRACT

Background: Accurate anthropometric measurement is fundamental for valid Body Mass Index (BMI) classification. In resource-variable settings like Nigeria, both analog and digital tools are used, yet there is a paucity of empirical evidence comparing their accuracy and reliability, particularly among university students. Methods: This cross-sectional comparative study recruited 475 undergraduates (age 18-30 years) via convenience sampling. Height and weight were measured using calibrated analog (SECA mechanical scale and stadiometer) and digital (Beurer BF 700) devices in a randomized order. Descriptive statistics, paired-sample t-tests, Pearson’s correlation, and Cronbach’s alpha were used for analysis. Results: Analog devices recorded a significantly greater mean height (169.16 ± 7.85 cm) than digital devices (167.43 ± 8.34 cm). Conversely, digital devices yielded significantly higher mean weight (64.82 ± 13.98 vs. 63.89 ± 14.01 kg) and BMI (23.12 ± 4.70 vs. 22.25 ± 4.36 kg/m²) (p < 0.001 for all comparisons). Despite these systematic differences, a very strong positive correlation (r = 0.94) and excellent reliability (Cronbach’s α = 0.97) were observed between the devices. Conclusion: While both device types are highly reliable and yield strongly correlated values, they produce statistically significant systematic differences in measurements. Therefore, analog and digital tools should not be used interchangeably for clinical or research purposes without adjustment, as the observed systematic differences in measurements could lead to BMI misclassification, affect the accuracy of population health screening, and undermine the standardization of research data. Digital devices are recommended for enhanced precision, whereas analog tools remain viable in low-resource settings provided rigorous calibration and standardized protocols are maintained.

Keywords: Body mass index (BMI); Analog BMI measurement tools; Digital BMI devices; Measurement, Accuracy, Reliability, students, Nigeria

# INTRODUCTION

One of the most commonly used anthropometric indicators for the assessment of body composition, nutritional status, and cardiometabolic risk in population studies is body mass index (BMI). It is commonly used in clinical and non-clinical settings due to its low cost, applicability and simplicity. It was initially developed in the 19th century by Adolphe Quetelet, and was standardized for global health use following the works of Keys and colleagues in 1972 (Keys et al., 1972). It is now widely adopted and it plays an important role in categorizing individuals into BMI groups, and the World Health Organization (WHO) classified adult weight status of BMI to: underweight (<18.5), normal (18.5 – 24.9), overweight (25.0 – 29.9), and obese (≥30.0) (WHO, 2022). While BMI remains relevant for its cost effectiveness and simplicity, it comes with some shortfalls as it cannot directly measure muscle mass, bone density, or fat distribution (Nuttall, 2015).

There has been a rising case of overweight and obesity among undergraduate students in Nigerian Universities and this is largely due to poor dietary choices where students consume more of fast food, sugary drinks, and snacks. Adding to this is a shift in their sedentary lifestyles where they focus more on screens (mobile devices) and do less walking, all of which is fueled by university independence and urbanization (Akindele et al., 2016; 6.	Abubaka et al., 2024). These lifestyle increases the vulnerability of undergraduate students to non-communicable diseases such as diabetes and hypertension (WHO, 2022).

Ability to accurately measure BMI values of students and counsel them for a change is dietary and sedentary lifestyle depend on the reliability of precise measurements tools for height and weight used to estimate BMI. One of the most widely used tools for measurement of height and weight for BMI in Nigeria are manual weighing scales and stadiometers, and this is because they are affordable and accessible (Stommel & Schoenborn, 2009). Digital BMI machines, equipped with load cells and ultrasonic sensors, automate measurements and reduce user error; however, their accuracy varies depending on environmental conditions in which they are used and calibration approach. Since a small difference in BMI values can result in misdiagnosis, improper health referrals, and misclassification (Nuttall, 2015), it is very essential to validate the accuracy of measurement tools before been used (Yorkin et al., 2013).

Some studies have shown that digital devices outperform analog instruments in terms of reproducibility and accuracy (Asa Berger, 2023; Bachin, 2024; Cerghizan et al., 2025; Koh et al., 2025). However, there is paucity of studies comparing analog and digital devices for BMI assessment, especially among undergraduate students in Nigeria. Hence, this study provides a direct comparative assessment of analog and digital devices for determination of BMI among undergraduate students in Nigerian universities, and offering an evidence-based insights for standardizing anthropometric measurement practices in academic institutions. To address the paucity of studies comparing analog and digital BMI assessment tools among Nigerian undergraduates, this study conducted a comparative assessment to: (1) measure and compare weight, height, and BMI using both device types; (2) evaluate the agreement, correlation, and systematic differences between the resulting BMI values; (3) assess the relative reliability of each method; and (4) derive evidence-based implications for standardizing anthropometric practices in resource-variable academic settings.

# MATERIALS AND METHODS

## Study design

This study employed a cross-sectional comparative design to evaluate the agreement between analog and digital tools for BMI assessment. Measurements of weight, height, and calculated BMI were obtained from each participant using both device types in a single session, following a standardized protocol to facilitate a direct within-subjects comparison.

## Study population

The study participants were 475 undergraduate students (aged 18–30 years) recruited from the Department of Anatomy and Physiology of the Faculty of Basic Clinical Science at the University of Port Harcourt, Nigeria. This department was selected as a single, convenient sampling frame due to logistical constraints and the researchers' affiliation, which facilitated access and participant recruitment.

## Recruitment and sampling techniques

A convenience sampling approach was employed for participant recruitment. Participants were recruited on-site within the Department of Anatomy and Physiology. Research assistants approached students in common areas such as the department lobby, study halls, and after selected lectures. The study's purpose and procedures were explained verbally, and interested students were screened for eligibility (age 18–30 years, having completed at least two semesters). Eligible individuals who provided written informed consent were enrolled and scheduled for measurement sessions.
While efforts were made to recruit across different academic levels within the department, participation was ultimately voluntary and dependent on students' presence and willingness during the data collection period.

## Instruments

Analog measurement tools
Weight was measured using a mechanical spring balance (Model: SECA 761, SECA GmbH & Co. KG, Hamburg, Germany) with a capacity of 150 kg and a graduation of 0.5 kg. Height was measured with a portable, wall-mounted mechanical stadiometer (Model: SECA 213, SECA GmbH & Co. KG) featuring a measuring range of 20–220 cm and a precision of 0.1 cm.

Digital measurement tool
An integrated, standalone digital BMI machine (Model: Beurer BF 700, Beurer GmbH, Ulm, Germany) was employed. This device incorporates a built-in, high-precision load-cell sensor for weight measurement (capacity: 180 kg, resolution: 0.1 kg) and an ultrasonic sensor for automatic height measurement (range: 60–210 cm, resolution: 0.1 cm). The device automatically calculates and displays the BMI on an integrated digital screen.

Calibration and environmental conditions
All devices were calibrated immediately prior to data collection and once weekly thereafter, adhering strictly to the manufacturers' instructions. The analog scale was zeroed before each measurement session, and the digital device underwent its internal electronic calibration routine. A standard 20 kg calibration weight was used for periodic verification of both scales, selected because it falls within the common measurement range for the study population, thereby ensuring accuracy across a critical segment of the expected weight values. All measurements were conducted in a dedicated, well-lit room with ambient temperature maintained between 22–25°C. Every device was placed on a firm, level, non-carpeted concrete floor to ensure stability and measurement accuracy.

## Measurement procedure

Participant preparation
Participants were instructed to wear light, standard clothing (such as t-shirt and shorts or light trousers), to remove their shoes, socks, heavy outerwear, and any items from their pockets, and to empty their bladder prior to measurement to minimize variability.

Measurement protocol
To control for order bias, the sequence of device use (analog or digital first) was randomized for each participant using a computer-generated random number list. The same trained research assistant performed all measurements for a given participant to ensure consistency in technique; however, the assistant was not blinded to the device readouts. For the analog stadiometer, participants stood erect with their back against the vertical rod, heels together, and head positioned in the Frankfort horizontal plane. The movable headboard was gently lowered to contact the vertex of the head. For the digital device, participants stood centered on the scale platform, and the ultrasonic sensor automatically detected height once they remained still. For both devices, participants stood calmly in the centre of the scale platform with arms relaxed at their sides. For the analog scale, the reading was recorded once the needle stabilized.

Each parameter (height and weight) was measured twice consecutively per device. If the two readings differed by more than 0.5 cm for height or 0.2 kg for weight, a third measurement was taken, and the two closest values were averaged. The mean value was used for all subsequent analyses.

Data recording and calculation
For the analog tools, the mean measured height (converted to meters) and weight (in kilograms) were manually recorded. BMI was subsequently calculated using the standard formula:

BMI (kg/m2) = Weight (kg)Height (m2)

For the digital device, the weight, height, and automatically calculated BMI were recorded directly from the digital display following the two measurement cycles. All data were entered immediately into a secure, password-protected electronic spreadsheet.

## Data analysis

The collected data were analyzed to compare measurements from the analog and digital devices. Descriptive statistics, including means and standard deviations (SDs), were calculated for height, weight, and BMI from each device type. A paired-sample t-test was used to determine if there were statistically significant differences between the mean values obtained from the two measurement systems. The Pearson correlation coefficient (r) was computed to assess the strength and direction of the linear relationship between the BMI values derived from the analog and digital tools. Finally, the internal consistency reliability of the two measurement methods was evaluated using Cronbach’s alpha (α). The threshold for statistical significance was set at p < 0.05 for all tests. Analyses were performed using SPSS software (version 25; IBM Corp., Armonk, NY, USA).

# RESULTS

## Comparative descriptive statistics of anthropometric parameters

The comparative analysis of anthropometric parameters revealed notable differences between measurement devices: analog tools recorded a greater mean height (169.16 ± 7.85 cm) compared to digital devices (167.43 ± 8.34 cm), while digital tools produced higher mean values for both weight (64.82 ± 13.98 kg vs. 63.89 ± 14.01 kg) and BMI (23.12 ± 4.70 kg/m² vs. 22.25 ± 4.36 kg/m²). Further examination of the distribution showed that digital measurements exhibited slightly wider variability across percentiles, particularly evident in the elevated 75th, 90th, and 95th percentile values for digital BMI (25.11, 29.63, and 32.22 kg/m², respectively) compared to analog (24.28, 28.03, and 30.10 kg/m²), suggesting a systematic upward shift in classification, especially in the higher BMI ranges.

Table 1: Descriptive statistics and percentiles of the Analog and digital anthropometric measures used in the study

| Parameters | Mean | SD | Min | Max | Percentiles |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  | 25% | 50% | 75% | 90% | 95% |
| Height (Analog) (cm) | 169.16 | 7.854 | 149 | 190 | 163 | 168.5 | 174 | 180.04 | 183.34 |
| Weight (Analog)(kg) | 63.8894 | 14.007 | 40 | 137 | 54 | 62.0 | 70.5 | 80.5 | 89.0 |
| BMI (Analog) (kg/m2) | 22.2547 | 4.359 | 14.43 | 45.25 | 19.170 | 21.249 | 24.281 | 28.034 | 30.096 |
| Height (digital) (cm) | 167.43 | 8.336 | 135 | 196 | 161.5 | 166.6 | 172.8 | 178.4 | 182.3 |
| Weight (digital) (kg) | 64.82 | 13.982 | 40 | 136 | 55.2 | 62.8 | 71.7 | 82.1 | 89.1 |
| BMI (digital) (kg/m2) | 23.12 | 4.704 | 15 | 46 | 19.96 | 21.94 | 25.11 | 29.63 | 32.22 |

Note: BMI: Body mass index

## Reliability and Agreement Between Analog and Digital BMI Measurements

The reliability analysis, presented in Table 2, indicates a very high degree of internal consistency between the analog and digital devices for BMI assessment. This is evidenced by an exceptionally strong positive correlation (Pearson's R = 0.94, r² = 0.89) and an excellent Cronbach's alpha coefficient (α = 0.97). These statistical results demonstrate that measurements from the two tools are highly reliable and produce consistently related values.

Table 2: Reliability and consistency of the Analog and digital BMI machine

|  | Mean ± SD | R (r2) | Cronbach alpha | Scale |
| --- | --- | --- | --- | --- |
| BMI (Analog) | 22.25 ± 4.35 | 0.94 (0.89) | 0.97 | Very reliable |
| BMI (digital) | 23.10 ± 4.70 |  |  |  |

## Statistical Comparison of anthropometric measures between device types

A paired-sample t-test revealed statistically significant differences between the analog and digital devices for all three anthropometric parameters (Table 3). While the analog device recorded a significantly greater mean height (169.16 ± 7.85 cm) than the digital device (167.29 ± 8.05 cm) [t = 10.937, p < 0.001], the digital device yielded significantly higher mean values for both weight (64.76 ± 13.93 kg vs. 63.88 ± 14.00 kg; t = -9.056, p < 0.001) and BMI (23.11 ± 4.70 kg/m² vs. 22.25 ± 4.35 kg/m²; t = -12.011, p < 0.001).

Table 3: Comparison of anthropometric measures using Analog device and digital devices

| Parameters | Analog | Digital | t-test | p-value | Inference |
| --- | --- | --- | --- | --- | --- |
| Height (cm) | 169.16 ±7.85 | 167.29 ± 8.05 | 10.937 | <0.001 | Significant |
| Weight (kg) | 63.88 ±14.00 | 64.76 ±13.93 | -9.056 | <0.001 | Significant |
| BMI (kg/m2) | 22.25 ± 4.35 | 23.11 ± 4.70 | -12.011 | <0.001 | Significant |

# DISCUSSION

This study provides a comparative analysis of analog and digital devices for anthropometric measurement among Nigerian university students. The key findings indicate that while both device types demonstrate excellent mutual reliability and a strong linear association, they produce statistically significant and systematic differences in mean values for height, weight, and consequently, BMI. Analog devices recorded a higher mean height, whereas digital devices yielded higher mean weight and BMI.

Our study revealed a significant difference in mean height, with analog devices recording higher values. This overestimation is consistent with known limitations of manual stadiometers, primarily attributed to parallax error and inter-operator variability in aligning the headpiece and reading the scale (Habicht 1982; Ulijaszek & Kerr, 1999). In contrast, digital ultrasonic sensors automate this process, potentially reducing this observer-dependent bias. However, it is important to note that digital devices may be sensitive to posture and hair styling, which could contribute to the lower mean observed (Stomfai et al., 2011). This discrepancy highlights that even simple height measurement is subject to tool-specific bias, which directly propagates into BMI calculation errors.

Conversely, digital devices produced significantly higher mean. This aligns with evidence suggesting that digital load-cell scales offer greater precision and are less susceptible to mechanical drift and reading errors compared to spring-based analog scales (Sebo et al., 2015; Nuttall, 2015). The analog scale’s slight underestimation of weight could be due to friction within the spring mechanism or less precise graduation interpretation. Our results corroborate studies that recommend digital scales for improved accuracy in epidemiological research where precise weight measurement is crucial (Gorber et al., 2007).

The most consequential result is the significantly higher mean BMI calculated from digital measurements. This systematic upward shift, particularly evident in the higher percentiles, suggests a tangible risk of misclassification when switching between device types. For instance, an individual measured with an analog device might be classified as "normal weight," whereas the same individual measured digitally could cross the threshold into the "overweight" category (BMI ≥ 25.0 kg/m²). This has direct implications for clinical assessment, public health statistics, and individual health counselling. Our findings are consistent with methodological studies emphasizing that even highly correlated tools can exhibit clinically relevant mean differences (Shields et al., 2008; Krul et al., 2011).

The excellent reliability and strong correlation between devices indicate they rank individuals consistently on the BMI continuum. This is reassuring for longitudinal studies using the same device type. However, as Bland and Altman have established, high correlation does not imply agreement (Bland & Altman, 1986). Our significant paired t-test results for all parameters confirm the presence of systematic bias, reinforcing that these measurements cannot be pooled or used interchangeably without adjustment. This distinction between reliability and agreement is paramount for correct data interpretation.

Study limitations and strengths
Several limitations must be acknowledged when interpreting these findings. First, the use of a convenience sample from a single academic department limits the generalizability of our results to the wider Nigerian student population or other demographic groups. Second, the absence of a gold-standard criterion measure for body composition or a calibrated high-precision scale prevents us from concluding which device type is more accurate in an absolute sense; we can only report on their relative differences and consistency. Third, while environmental conditions were controlled, unmeasured factors such as subtle variations in floor levelness or participant hydration status could have introduced noise. The strength of this study lies in its direct, within-subject comparative design, a relatively large sample size for a device-comparison study, and the use of standardized procedures, which bolsters the internal validity of the observed differences.

## Recommendations

Based on the evidence presented, we recommend that digital BMI devices be prioritized for research and clinical screening where resources permit, as their automated functions reduce observer bias and enhance measurement precision, which is particularly crucial for multi-center studies and national surveillance programs requiring standardization. In low-resource settings reliant on analog tools, rigorous and frequent calibration against standard weights, coupled with comprehensive staff training on standardized techniques (such as Frankfort plane alignment, scale zeroing), is essential to minimize systematic error. Furthermore, researchers and practitioners must explicitly report the type of device used when presenting anthropometric data and exercise caution when comparing datasets derived from different device types or interpreting longitudinal data following a change in measurement tool. Future studies should incorporate a gold-standard validation method to establish the true accuracy of both analog and digital devices in this context and explore cost-effective calibration protocols and training interventions for analog tools in field settings.

## Conclusion

This study demonstrates that the choice of anthropometric tool has a statistically and clinically significant impact on BMI assessment; therefore, moving toward standardized, well-calibrated digital measurement protocols, while implementing strict quality control for analog tools, is a necessary step for improving the accuracy and comparability of BMI data in Nigerian academic and similar resource-variable settings.

ETHICAL APPROVAL

Ethical clearance was obtained from the University of Port Harcourt Research Ethics Committee. Informed consent was obtained from all participants.

Disclaimer (Artificial intelligence)

Option 1:

Author(s) hereby declare that NO generative AI technologies such as Large Language Models (ChatGPT, COPILOT, etc.) and text-to-image generators have been used during the writing or editing of this manuscript.

Option 2:

Author(s) hereby declare that generative AI technologies such as Large Language Models, etc. have been used during the writing or editing of manuscripts. This explanation will include the name, version, model, and source of the generative AI technology and as well as all input prompts provided to the generative AI technology

Details of the AI usage are given below:

1.

2.

3.

# REFERENCES

- Keys, A., Fidanza, F., Karvonen, M. J., Kimura, N., & Taylor, H. L. (1972). Indices of relative weight and obesity. Journal of Chronic Diseases, 25(6–7), 329–343. https://doi.org/10.1016/0021-9681(72)90027-6

- World Health Organization. (2022). BMI classification and measurement standards. WHO. https://www.who.int/data/gho/data/themes/topics/topic-details/GHO/body-mass-index

- Nuttall F. Q. (2015). Body Mass Index: Obesity, BMI, and Health: A Critical Review. Nutrition today, 50(3), 117–128. https://doi.org/10.1097/NT.0000000000000092

- Akindele, M. O., Phillips, J. S., & Igumbor, E. U. (2016). The Relationship Between Body Fat Percentage and Body Mass Index in Overweight and Obese Individuals in an Urban African Setting. Journal of public health in Africa, 7(1), 515. https://doi.org/10.4081/jphia.2016.515

- Abubakar, H. A., Shahril, M. R., & Mat, S. (2024). Nutritional status and dietary intake among Nigerian adolescent: a systematic review. BMC public health, 24(1), 1764. https://doi.org/10.1186/s12889-024-19219-w

- Stommel, M., & Schoenborn, C. A. (2009). Accuracy and usefulness of BMI measures based on self-reported weight and height: findings from the NHANES & NHIS 2001-2006. BMC public health, 9, 421. https://doi.org/10.1186/1471-2458-9-421

- Yorkin, M., Spaccarotella, K., Martin-Biggers, J., Quick, V., & Byrd-Bredbenner, C. (2013). Accuracy and consistency of weights provided by home bathroom scales. BMC public health, 13, 1194. https://doi.org/10.1186/1471-2458-13-1194

- Asa Berger, A. (2023). Analog and Digital. Journal of Cyberspace Studies, 7(2), 105–114. https://doi.org/10.22059/jcss.2023.94683

- Bachin, D. (2024). Enhancing reproducibility in research through digitalization. SciSure. https://www.scisure.com/blog/enhancing-reproducibility-in-research-through-digitalization

- Cerghizan, D., Jánosi, K. M., Farcas, A., Bojan, M. M., Muntean, M. H., Nechiti, A. A. M., Mureșan, I. É., Pop, S. I., & Marada, G. (2025). Accuracy of Analog and Digital Full-Arch Mandibular Impressions: In Vitro and In Vivo Evaluation. Diagnostics (Basel, Switzerland), 15(16), 2077. https://doi.org/10.3390/diagnostics15162077

- Koh, Y.-S., Naidoo, N., & Petridis, H. (2025). Comparisons of analog and digital methods to produce an accurate trial restoration. The Journal of Prosthetic Dentistry, 133(2), 513–522. https://doi.org/10.1016/j.prosdent.2024.02.012

- Habicht, J. P., Meyers, L. D., & Brownie, C. (1982). Indicators for identifying and counting the improperly nourished. The American journal of clinical nutrition, 35(5 Suppl), 1241–1254. https://doi.org/10.1093/ajcn/35.5.1241

- Ulijaszek, S. J., & Kerr, D. A. (1999). Anthropometric measurement error and the assessment of nutritional status. The British journal of nutrition, 82(3), 165–177. https://doi.org/10.1017/s0007114599001348

- Stomfai, S., Ahrens, W., Bammann, K., Kovács, E., Mårild, S., Michels, N., Moreno, L. A., Pohlabeln, H., Siani, A., Tornaritis, M., Veidebaum, T., Molnár, D., & IDEFICS Consortium (2011). Intra- and inter-observer reliability in anthropometric measurements in children. International journal of obesity (2005), 35 Suppl 1, S45–S51. https://doi.org/10.1038/ijo.2011.34

- Sebo, P., Haller, D., Pechère-Bertschi, A., Bovier, P., & Herrmann, F. (2015). Accuracy of doctors' anthropometric measurements in general practice. Swiss medical weekly, 145, w14115. https://doi.org/10.4414/smw.2015.14115

- Connor Gorber, S., Tremblay, M., Moher, D., & Gorber, B. (2007). A comparison of direct vs. self-report measures for assessing height, weight and body mass index: a systematic review. Obesity reviews : an official journal of the International Association for the Study of Obesity, 8(4), 307–326. https://doi.org/10.1111/j.1467-789X.2007.00347.x

- Krul, A. J., Daanen, H. A., & Choi, H. (2011). Self-reported and measured weight, height and body mass index (BMI) in Italy, the Netherlands and North America. European journal of public health, 21(4), 414–419. https://doi.org/10.1093/eurpub/ckp228

- Shields, M., Gorber, S. C., & Tremblay, M. S. (2008). Effects of measurement on obesity and morbidity. Health reports, 19(2), 77–84. https://pubmed.ncbi.nlm.nih.gov/18642521/

- Bland, J. M., & Altman, D. G. (1986). Statistical methods for assessing agreement between two methods of clinical measurement. Lancet. https://doi.org/10.1016/S0140-6736(86)90837-8