



Production of Shape Memory and Biodegradable Fe-Mn-Si-Ti-Mg-Zn-C Alloy Foam for Temporary Implant Applications
Abstract

In this study, highly porous shape memory, nonmagnetic and biodegradable Fe alloy foam was fabricated for temporary implant and scaffold applications. Mg alloys biodegrade too fast with H2 evolution. Zn alloys show biodegradation rates in the middle of Mg and Fe alloys, but the Zn alloys are very brittle. Zn also shows low strength and low plastic deformation. Fe has advantages over Mg and Zn alloys for the stent applications in terms of mechanical properties. Fe alloys also show radio-opaqueness, which is important in the coronary stents. Fe alloy specimens with open porous structure were fabricated by space holder method. Carbamide was used as a space holder. Biodegradation rate was investigated by using weight loss and metal ion release measurements. Corrosion performance was tested in simulated body fluid. Cytotoxicity of the foams was studied by neutral red uptake assay. Microstructure was investigated by using optical microscope and scanning electron microscope. Fe-4Mn-4Si-2Ti-5Mg-5Zn-0.5C alloy does not have a cytotoxic potential on the cells. Fe-4Mn-4Si-2Ti-5Mg-5Zn-0.5C alloy has showed shape memory effect. Fe-4Mn-4Si-2Ti-5Mg-5Zn-0.5C alloy does not have a cytotoxic potential on the cells according to neutral red uptake (NRU) assay. Fe-4Mn-4Si-2Ti-5Mg-5Zn-0.5C alloy has showed suitable biodegradation rate for temporary implant applications. 
Keywords: Shape memory, Metal foam; Biodegradation; Fe alloy; Powder metallurgy; 
1.  Introduction

Highly porous biodegradable scaffolds are used in the tissue engineering. In general, biodegradable scaffold provides mechanical support for the seeded living cells and determines the shape of the tissues. Scaffold allows the transportation of the body fluids by the open pores and provides tissue growth. Scaffold materials should have a definite optimum biodegradation rate. Meanwhile, biodegradable metals can be used in temporary orthopaedic implant applications. Complete dissolution of the biodegradable implants assists tissue healing without any residues. Permanent metallic orthopaedic implants show problems such as inflammation, thrombosis, stress shielding. Moreover, permanent orthopaedic metal implants need a secondary surgical operation [1-3]. 

Traditional paradigm of metallic biomaterials requires metals with high electrochemical corrosion resistance. Biodegradable metals were changed this paradigm. Biodegradable metals are metals expected to corrode gradually with a suitable host response. Biodegradable metals can be classified as biodegradable pure metals, biodegradable alloys, and biodegradable metal matrix composites. Polymer biomaterials have low strength, low wear resistance and release toxic chemical reagents, while ceramic biomaterials are very brittle. Biodegradable metals are metals expected to corrode gradually with a suitable host response. Magnesium (Mg), iron (Fe) and zinc (Zn) are biodegradable metals [1-3]. 
Mg content is about 25-40 gr in the bones. Mg is also important for enzyme reactions. In general, the high electrochemical corrosion rate of the Mg alloys is attributed to its low standard electrode potential. In addition, corrosion products can be metabolized or digested. Moreover, Mg alloys can stimulate bone formation, which is important in the bone fracture healing. Mg alloys biodegrade too fast with the H2 evolution. Mechanical properties of the Mg alloys are very low for the implant applications but eliminates the possible toxic effects of the alloying elements. 

Zn alloys show biodegradation rates in the middle of Mg and Fe, and their biodegradation by-products are biodegradable. Electrochemical corrosion of the Zn does not produce hydrogen due to the high hydrogen over-potential of the Zn. Daily intake of the Zn is about 40 mg. Zn is very brittle. Zn shows low strength and low plastic deformation. In addition, low fatigue properties, recrystallization at low temperatures, and susceptibility to aging may lead to failure of the implants. 
Fe alloys are biodegradable and non-toxic. Fe alloys also show radio-opaqueness so that markers to make the coronary stent visible by fluoroscopy is not required [4-7]. Pure Fe has advantages over Mg alloys and Zn alloys for the stent applications in terms of mechanical properties. One method to fasten the biodegradation rate is addition of some biocompatible elements with lower standard electrode potential into the Fe. Second method is addition of alloying elements into the Fe in order to produce phases which produce galvanic corrosion and, increase the corrosion rate [8-11]. 
Open porous structure is important in the implants for the tissue ingrowth. Open-cell foams can be manufactured by powder metallurgy method [11-15]. There are no solubility limitations in powder metallurgy as compared with casting. Repeatability of the powder metallurgy method is also superior [9-18]. Sadighikia et al. [8] have manufactured Zn foams with open-cellular porous microstructure by using space holder technique. They used carbamide as a pore former. Francis et al. [9] studied Fe alloys for stent applications. Fe alloy based implants shows strength, high corrosion rate and a suitable biocompatibility. For their suitable biocompatibility and suitable interaction with the surrounding living tissue, Fe alloy stents may biodegrade slowly while retaining their mechanical integrity. Potential of Fe alloys was attributed to the reduction of side effects such as inflammation and restenosis. Research on alloy development should be continued with biodegradation. Orinakova et al. [12] produced porous d iron alloy foams for orthopaedic implant applications. Highly porous iron foams were fabricated by replication method by using foamed polyurethane. Capek et al. [13] produced highly porous Fe foams with 30-80 % porosity by powder metallurgy. Ammonium bicarbonate was used as a pore former. 
Shape memory alloys show shape recovery to their pre-deformed first shape by heating or by removal of stress. An alloy must have following abilities for shape memory behaviour; (1) thermoelastic behaviour, (2) twinning, (3) transformation to ordered austenite structure. Shape memory effect is due to diffusionless solid thermoelastic-martensitic phase transformation. Plastic deformation must be by twinning not by slip of dislocations, and martensitic phase transformations must be thermoelastic, and there must be orientation [16, 17].
In the present study, Fe alloy foams were fabricated for temporary implant applications by the space holder method. Novelties are (1) shape memory, nonferromagnetic and biodegradable foam production and (2) lower alloying element contents in order to obtain nonferromagnetic structure, shape memory and biodegradation. There are studies on the shape memory FeMnSi alloys and biodegradable FeMnSi alloys. There is no study on the porous shape memory and biodegradable FeMnSi alloy. In literature, there are studies on the biodegradable Fe alloys, which are includes high amount of Mn. There is no study on the FeMnSiTiMZnC alloy, which includes lower amounts of alloying elements. In the present study, ferromagnetic properties were eliminated by lower alloying element addition. Mn addition produces austenite which eliminates the ferromagnetic properties. Wang et al. [18] studied biodegradation properties of the Fe, Fe35Mn alloy, and (Fe35Mn)5Ag alloy, which were produced by powder metallurgy. Hermawan et al. [19] studied Fe-Mn based alloys and concluded that the biodegradation rate of the Fe-Mn alloys lowered with increasing of the Mn content from the 20 % to 35 %.
2. Experimental
2.1. Alloy Production 
Fe alloy foams were fabricated by using metal powders. The average particle size of the powders was 44 μm. At the beginning, the metal powder mixtures were ball-milled (mechanical alloying) with the zirconia balls (6 mm diameter) for about 20 hours with a rotational speed of 400 rpm. Metal powder to zirconia ball ratio was about ten to one. Highly porous specimens were fabricated by powder metallurgy based space holder method. In the powder metallurgy method, carbamide powder in the range of 710-1000 μm was employed as a pore former (space holder). Mean particle size of the carbamide was about 800-900 μm. Mixtures were pressed at 190-200 MPa into the cylindrical specimens with diameters of 20 mm and 12 mm and heights of 15-17 mm. Carbamide was removed in the water and then the samples were sintered for 1 hour under vacuum. Sintering temperature was 1150 ºC. 

Schaeffler diagram can be used in order to predict the phases of the Fe, by using Niequivalent and Crequivalent, which are given in equations (1) and (2). Implant materials must be magnetic resonance imaging (MRI) compatible. Nonferromagnetic austenite phase must be obtained in the Fe. In the present study, nonferromagnetic austenite Fe4Mn4Si2Ti5Mg5Zn0.5C alloy with Niequivalent of 19% and Crequivalent of 11% was produced [20].
Crequivalent= Cr+2Si+1.5Mo+5V+5.5Al+1.75Nb+1.5Ti+0.75W 


(1)
Niequivalent= Ni+Co+0.5Mn+0.3Cu+25N+30C 




(2)
2.2. Shape Memory Training

Free shape memory behaviour (amount of shape recovery) of the alloy was investigated by measuring the height 12 mm diameter and 16 mm height cylindirical specimens under uniaxial compression. Initially, cylindirical specimen was heated to 900 °C for 10 minutes and then quenched. Secondly, height of the cylindrical specimen (h0) was measured. Next, a uniaxial compression stress was applied on the cylindrical shaped specimen until predetermined 2 % strain and height was measured (h1). Lastly, after removal of the stress, the specimen was heated to 900 °C (Af) and shape recovery was determined by measuring the last height (h2) of the specimen. Shape recovery (r) was determined by using the following equation (3) [21].


[image: image1.wmf]100

1

(%)

r 

0

1

0

2

´

÷

÷

ø

ö

ç

ç

è

æ

-

-

-

=

h

h

h

h









 (3)

Thermal cycling included exposing the specimen to repeating thermal cycles between 900 °C and 80 °C under of a fixed external bias stress (25 MPa). After heating step at 900 °C, the tube was removed out from the furnace. A cooling fan was used for faster cooling. Subsequently, the tube was slide inside the furnace until the hot zone.

2.3. Electrochemical Corrosion and Biodegradation Tests
Simulated body fluid was prepared from chemicals according to the literature [22]. The amounts of the chemicals (in g/L) were 8.0 NaCl, 0.3 CaCl2, 0.2 KCl, 0.3 MgCl2, 0.2 K2HPO4, 0.35 NaHCO3, 0.07 Na2SO4, 6.0 tris and 1.0 M HCl. The pH of the SBF solution was 6.60. Electrochemical corrosion tests were done using a potentiostat (Interface, Gamry). Initially, open-circuit potential was conducted. Tafel and linear polarization resistance tests were employed in order to determine the corrosion rates of the samples. Fe alloy specimens were dipped into the simulated body fluid (SBF) for biodegradation tests. Weight loss values (%) were determined by gravimetric method. After different periods, the samples were removed from the solution. Weight loss (%) was computed by weighting the dried samples.
2.4. Cytotoxicity Evaluation

In the present study, in vitro evaluation of the cytotoxicity of the Fe alloy foams were studied by extraction based 3T3 neutral red uptake (NRU) assay. The 3T3 neutral red uptake (NRU) assay is based on the ability of living cells to uptake supravital neutral red dye, which can penetrate the cell membrane by nonionic diffusion. Neutral red (NR) is a weak cationic dye. Living cells could be distinguished from the dead cells. The neutral red uptake assay provides a quantitative estimation of the number of viable cells in a culture. Positive (sodium lureth sulfate) and negative (polypropylene) samples were prepared to verify the test system. The sterilized Fe alloy test substances were exposed to immortalised mouse fibroblast Balb/c 3T3 cell line. Absorbance measurements were performed at 540 nm wavelengths in order to determine the viability/survival. Cytotoxicity can be expressed as reduction of the uptake of neutral red. Dulbecco’s Modification of Eagle’s Medium (DMEM) cell culture media, bovine calf serum (BCS), trypsin/EDTA were the chemicals. Serum culture medium (SCM) was used as extraction agent, as the SCM supports extraction of polar and non-polar substances and cell proliferation. Cell culture was removed from culture flasks by enzymatic digestion (trypsin/EDTA) and then the cell suspension was centrifuged. The cells then were resuspended and seeded at 104 cell per well in DMEM in a 96-well plate. After incubation the culture was aspirated. After incubation, the culture medium was removed and the cells were washed with phosphate buffered saline (PBS), The test system is considered suitable if the viability for the negative control specimen is >70 % of blank.
3. Results and Discussion
3.1 Microstructure Characterization 

Biodegradable and biocompatible highly porous specimens were fabricated for the temporary implant applications. Figure 1 shows the optic microscope picture from the cell-wall. Microstructure consists of austenite main phase (bright, white areas) and ferrite phase (brown areas). There are some micropores (black areas), which are necessary for transportation of the body fluids. Magnetic flux density was 0.97 mT, which confirms that the alloy is nonferromagnetic. Meanwhile, specimen was not being pulled by a permanent magnet.
[image: image2.jpg]


   
Figure 1. Optic microscope picture from the cell-wall of the porous specimen
Figure 2 shows the photographs of the (a) sintered porous Fe alloy specimen, and (b) cross-section of the sintered highly porous Fe alloy specimen. There is no micro-cracks and macro-cracks on the surface of the sintered specimen.
[image: image3.jpg]



Figure 2. Photographs of the a) porous Fe alloy and b) cross-section of porous Fe alloy
Figure 3 illustrates the corrosion rates of the Fe alloys. Corrosion rates of the specimens were high. Fe-Mn-Si-Ti-Mg-Zn-C alloy also show high corrosion rate. Corrosion rate of the Fe-Ti specimen was lower than the pure Fe specimen. Effect of Mn and Si addition was neutral.

[image: image4.jpg]Corrosion Rate (mpy)

0,4

0,3 1

ol

'

A

'

Alloy

ce?®

N
S

<@

N\g’L“G




Figure 3. Corrosion rates values of the Fe alloys
Figure 4 illustrates the weight change values of the alloys for 21 days. Zn and Mg additions were increased the metal ion release. Effect of Mn and Si addition was neutral. 

[image: image5.jpg]Weight Change (%)

20

15





Figure 4. Weight change values of the porous Fe alloys

Figure 5 illustrates the variation of the Young’s modulus with immersion time in the simulated body fluid. Young’s modulus values were deceased with immersion time.
[image: image6.jpg]Young's Modulus (GPa)

12

10 4

B 7 days
[ 14 days
[ 1 21 days

?ef‘;L(\ ?e’(Z“\ 6\\\(\ ?G’b‘s\ 6\\\9 \\\g«L(\O




Figure 5. Variation of the Young’s modulus of the porous Fe alloys with immersion time
3.2 Shape Memory Properties  
Free shape recovery was investigated by measuring the height of the specimens. Initially, cylindirical specimen was heated to 900 °C for 10 minutes and then quenched. Secondly, height of the quenched specimen was measured. Next, a uniaxial compression stress was applied on the cylindrical shaped specimen and the height was measured. Lastly, after removal of the stress, the specimen was heated to 900 °C (Af) and shape recovery was determined by measuring the last height. Thermal cycling includes exposing the specimen to repeating thermal cycles. Figure 6 shows the effect of thermal cycles on the shape memory behaviour (shape recovery) of the alloy. As seen from the figure, increasing number of thermal cycles increased the amount of shape recovery ability. Thermal cycling (training) affects phase transformation temperatures, and affect shape memory ability.
[image: image7.png]Recovery (%)

16

14 4

12 4

10 4

6 8 10 12
Thermal Cycle (number)

18




Figure 6. Effect of thermal cycles on the shape recovery of the alloy
3.3 Toxicity 

In the present study, in vitro evaluation of the cytotoxicity of the Fe alloy foams were studied by extraction based 3T3 neutral red uptake (NRU) assay. The 3T3 neutral red uptake (NRU) assay is based on the ability of living cells to uptake supravital neutral red dye, which can penetrate the cell membrane. The sterilized Fe alloy test substances were exposed to immortalised mouse fibroblast Balb/c 3T3 cell line. Viability of Fe-4Mn-4Si-2Ti-5Mg-5Zn-0.5C alloy was 73 %. This means that the alloy has no cytotoxic potential. Cytotoxic values indicated the potential applications of the Fe-4Mn-4Si-2Ti-5Mg-5Zn-0.5C alloy. Meanwhile, viability of the sintered pure Fe specimen was 79 %. 

4. Conclusions
In this study, highly porous, nonferromagnetic, biocompatible and biodegradable Fe alloy foams were fabricated for temporary implant and scaffold applications by powder metallurgy based space holder method.

The main conclusions of the present study are as follows:

(1) Fe-4Mn-4Si-2Ti-5Mg-5Zn-0.5C alloy has showed shape memory effect (shape recovery)

(2) Fe-4Mn-4Si-2Ti-5Mg-5Zn-0.5C alloy does not have a cytotoxic potential on the cells according to neutral red uptake (NRU) assay.

(3) Fe-4Mn-4Si-2Ti-5Mg-5Zn-0.5C alloy has showed suitable biodegradation rate for temporary biomedical implant applications. 

(4) Microstructure of the Fe-4Mn-4Si-2Ti-5Mg-5Zn-0.5C alloy consists of austenite main phase and some ferrite phase.

References
[1] L. Zhao, Z. Zhang, Y. Song, S. Liu, Y. Qi, X. Wang, Q. Wang, C. Cui. Materials and Design 108 (2016) 136-144.

[2] Z.S. Seyedraoufin, S. Mirdamadi. Journal of the Mechanical Behaviour of Biomedical Materials 21 (2013) 1-8.

[3] H. Li, Q. Peng, X. Li, K. Li, Z. Han, D. Fang. Materials and Design 58 (2014) 43-51.

[4] D. Vojtech, J. Kubasek, J. Serak, P. Novak. Acta Biomaterialia 7 (2011) 3515-3522.

[5] S. Zhao, J.M. Seitz, R. Eifler, H.J. Maier, R.J. Guillory II, E.J. Earley, A. Drelich, J. Goldman, J.W. Drelich. Materials Science and Engineering C 76 (2017) 301-312.

[6] P.K. Bowen, R.J. Guillory II, E.R. Shearier, J.M. Seitz, J. Drelich, M. Bocks, F. Zhao, J. Goldman. Materials Science and Engineering C 56 (2015) 467-472.

[7] A.J. Drelich, S. Zhao, R.J. Guillory II, J.W. Drelich, J. Goldman. Acta Biomaterialia 58 (2017) 539-549.

[8] S. Sadighikia, S. Abdolhosseinzadeh, H. Asgharzadeh. Powder Met. 58, 1 (2015), 61-66.
[9] A. Francis, Y. Yang, S. Virtanen, A.R. Boccaccini. J Mater Sci: Mater Med (2015) 26:138.

[10] H.O. Gulsoy, R.M. German. Scripta Materialia 58 (2008) 295-298.
[11] F. Kafkas, T. Ebel. Journal of Alloys and Compounds 617 (2014) 359-366.

[12] K. Berent, J. Pstrus, T. Gancarz. Journal of Materials Engineering and Performance (2016) 25:3375-3383.

[13] A. Kafri, S. Ovadia, J. Goldman, J. Drelich, E. Aghion. Metals, (2018) 8, 153.
[12] R. Orinakova, A. Orinak, L.M. Buckova, Maria Giretova, L. Medvecky, E. Labbanczova, M. Kupkova, M. Hrubovcakova, K. Kova. Int. J. Electrochem. Sci., 8 (2013) 12451-12465.

[13] J. Capek, D. Vojtech, A. Oborna. Materials & Design 83 (2015) 468-482.

[14] D. Yang, J. Chen, W. Chen, L. Wang, H. Wang, J. Jiang, A. Ma. Journal of Materials Science & Technology 33 (2017) 1141-1146.
[15] Y. H. Wang, L. R. Xiao, X. J. Zhao, W. Zhang, Y. F. Song, L. Guo and Y. Wang. Materials and Corrosion, 67, 3 (2016), 297-304.

[16] B. Liu, Y.F. Zheng, Liquan Ruan, Materials Letters 65 (2011) 540-543

[17] Yong Nie, Bo Yuan, Jiawei Liang, Tao Deng, Xiaofan Li, Peng Chen, Kai Zhang, Xiangyun Li, Kang Li, Huabei Peng, Shan Gong. Materials Science & Engineering A 908 (2024) 146725

[18] C. Wang, C. Tonna, D. Mei, J. Buhagiar, M.L. Zheludkevich, S.V. Lamaka. Bioactive Materials 7 (2022) 412-425.

[19] H. Hermawan, D. Dube, D. Mantovani, Adv. Mater. Res. 15–17 (2006) 107–112. 

[20] ASM Handbook Volume 1. Second Edition. ASM International, 1993.

[21] B. Liu, Y.F. Zheng, Liquan Ruan, Materials Letters 65 (2011) 540-543

[22] T. Kokubo, H. Takadama. Biomaterials, 27 (2007) 2907-2915.
11
9

_1795769528.unknown

