| Journal Name: | # Journal of Medicine and Health Research |
| --- | --- |
| Manuscript Number: | Ms_JOMAHR_14986 |
| Title of the Manuscript: | An Integrated Multi-Modal Wearable Sensor System for Real-Time Monitoring of Athlete Stress, Hydration, and Tendon Stiffness Analysis |
| Type of the Article | Research Article |

PART 1 (Importance of the manuscript)

|  | ## Comments of the Reviewers | Author’s Feedback |
| --- | --- | --- |
| Please write a few sentences regarding the importance of this manuscript for the scientific community. A minimum of 3-4 sentences may be required for this part. | This manuscript addresses a clinically relevant gap in sports medicine by proposing a compact Arduino-based wearable platform for simultaneous real-time monitoring of athlete stress, hydration, and tendon mechanics. The convergence of galvanic skin response, bioimpedance analysis, high-resolution IMUs, and strain sensors in a single microcontroller-based framework is a meaningful systems engineering contribution, as most prior wearable systems are limited to one or two physiological modalities. If rigorously validated, such a platform could support injury-prevention protocols in settings where laboratory-grade monitoring is unavailable. However, substantial revision is needed before the system's novelty and feasibility can be fully evaluated | ## Thank You reviewer for the valuable comments and positive evaluation of the manuscript’s structure, sustainability focus, and clarity of presentation |

## PART 2.1 (Objective Evaluation)

|  | ## Rating of the Reviewers | Author’s Feedback |
| --- | --- | --- |
| 1. Is the title clear and appropriate for the study? Rating Scale: 5 = Excellent 4 = Good 3 = Satisfactory 2 = Needs Improvement 1 = Poor N/A = Not Applicable | 4 | ## Thank You |
| ## 2. Is the abstract of the article comprehensive? Rating Scale: 5 = Excellent 4 = Good 3 = Satisfactory 2 = Needs Improvement 1 = Poor N/A = Not Applicable | 3 | ## Thank You |
| ## 3. Are the keywords appropriate and useful? Rating Scale: 5 = Excellent 4 = Good 3 = Satisfactory 2 = Needs Improvement 1 = Poor N/A = Not Applicable | 4 | ## Thank You |
| ## 4. Is the background information of the paper sufficient and well organized? Rating Scale: 5 = Excellent 4 = Good 3 = Satisfactory 2 = Needs Improvement 1 = Poor N/A = Not Applicable | 4 | ## Thank You |
| ## 5. Are the research objectives/hypotheses clearly stated? Rating Scale: 5 = Excellent 4 = Good 3 = Satisfactory 2 = Needs Improvement 1 = Poor N/A = Not Applicable | 3 | ## Thank You |
| ## 6. Is the literature review relevant and up to date? Rating Scale: 5 = Excellent 4 = Good 3 = Satisfactory 2 = Needs Improvement 1 = Poor N/A = Not Applicable | 3 | ## Thank You |
| ## 7. Is the research methodology appropriate for the study? Rating Scale: 5 = Excellent 4 = Good 3 = Satisfactory 2 = Needs Improvement 1 = Poor N/A = Not Applicable | 3 | ## Thank You |
| ## 8. Were ethical issues properly addressed (if applicable)? Rating Scale: 5 = Excellent 4 = Good 3 = Satisfactory 2 = Needs Improvement 1 = Poor N/A = Not Applicable | 2 | ## Thank You |
| 9. Are the results presented clearly? Rating Scale: 5 = Excellent 4 = Good 3 = Satisfactory 2 = Needs Improvement 1 = Poor N/A = Not Applicable | 2 | ## Thank You |
| 10. Are tables and figures clear, relevant, and necessary? Rating Scale: 5 = Excellent 4 = Good 3 = Satisfactory 2 = Needs Improvement 1 = Poor N/A = Not Applicable | 2 | ## Thank You |
| 11. Does the discussion relate findings to existing literature? Rating Scale: 5 = Excellent 4 = Good 3 = Satisfactory 2 = Needs Improvement 1 = Poor N/A = Not Applicable | 2 | ## Thank You |
| 12. Are the conclusions supported by the data? Rating Scale: 5 = Excellent 4 = Good 3 = Satisfactory 2 = Needs Improvement 1 = Poor N/A = Not Applicable | 2 | ## Thank You |
| 13. Are the limitations of the study discussed? Rating Scale: 5 = Excellent 4 = Good 3 = Satisfactory 2 = Needs Improvement 1 = Poor N/A = Not Applicable | 1 | ## Thank You |
| 14. Are the references relevant and sufficient (in number)? Rating Scale: 5 = Excellent 4 = Good 3 = Satisfactory 2 = Needs Improvement 1 = Poor N/A = Not Applicable | 4 | ## Thank You |
| 15. Is the manuscript written in clear and understandable language? Rating Scale: 5 = Excellent 4 = Good 3 = Satisfactory 2 = Needs Improvement 1 = Poor N/A = Not Applicable | 3 | ## Thank You |

## PART 2.2 (Subjective Evaluation)

|  | ## Reviewer’s comment | Author’s Feedback (It is mandatory that authors should write his/her feedback here) |
| --- | --- | --- |
| Is the title of the article suitable? If your answer is NO, please provide a brief, clear suggestion for improvement. | Partially suitable. The phrase "Tendon Stiffness Analysis" overstates the contribution, as the system measures kinematic/kinetic surrogates via IMU and strain sensors rather than direct mechanical stiffness. | ## Sincerely thank you reviewer for the positive comments |
| ## Is the abstract of the article comprehensive? s If your answer is NO, please provide a brief, clear suggestion for improvement. | yes | ## Sincerely thank you reviewer for the positive comments |
| ## Is the manuscript scientifically correct? If your answer is NO, please provide a brief, clear suggestion for improvement . | Partially. Three specific concerns arise: (1) Characterizing IMU/strain-sensor outputs as "tendon stiffness" is inaccurate - these are kinematic proxies, not direct stiffness measurements; the terminology should be corrected throughout. (2) The feasibility of an ATmega328P at 16 MHz performing simultaneous HRV analysis, BIA processing, and IMU fusion in real-time is not demonstrated or benchmarked - this is a non-trivial computational claim. (3) The conversion of raw bioimpedance values to hydration estimates requires a validated algorithm that is not described anywhere in the methodology. | ## Sincerely thank you reviewer for the positive comments |
| Are the references sufficient and recent? (YES or NO) If your answer is NO, please provide clear suggestion for improvement. | Yes. The 29 references span 2018-2025, with strong representation of 2024-2025 publications across all four major domains (stress, hydration, tendon mechanics, multi-modal wearables). Minor inconsistency: some journal names appear in bold in the reference list while others do not - formatting should be standardized. | ## Sincerely thank you reviewer for the positive comments |
| Are there ethical issues in this manuscript? (YES or NO) (If yes, kindly please write down the ethical issues here in details) | Yes. The manuscript describes validation trials involving human subjects (10-30 athletes) but contains no mention of: (1) institutional review board (IRB) or ethics committee approval; (2) informed consent procedures; (3) data privacy protections. This omission must be rectified before publication. The authors should add an explicit ethics statement in the Methods section with the relevant approval number and consent process. | ## Sincerely thank you reviewer for the positive comments |